## Assessment Summary
Unit of Assurance — Assessment Summary
Project Name | COU Name | COU Description | Profile | Device Class | Model Risk Level | Assurance Level | Standards Reference | Assessor Name | Assessment Date | Source Document | Has UQ?
Alpha-52 Cementless Acetabular Cup Structural Model | Preclinical screening of porous Ti-6Al-4V acetabular shell for primary hips under gait loading | FEA-based structural model used to evaluate interface micromotion and rim contact conditions for a 52 mm cementless acetabular cup under worst-case stance-phase loading. The model informs a go/no-go decision for design freeze and cadaveric testing. Applicability covers shell sizes 48–56 mm, primary hips, bone quality T-score −1 to −2, loads up to 1.5 kN and 30 N·m torque, and seating misalignment within ±2°. | Complete | Class III | MRL 3 | Medium | ASME-VV40-2018 | R. Patel, Simulation Group | Date of this assessment (YYYY-MM-DD) | report.md | Yes
Legend: |  | Required (Minimal) | Required (Complete) | Optional
Note on A3: [Extract] ¦ Confidence: 50%
Note on B3: [Extract] ¦ Confidence: 50%
Note on C3: [Extract] ¦ Confidence: 50%
Note on D3: [Extract] ¦ Confidence: 50%
Note on E3: [Extract] ¦ Confidence: 50%
Note on F3: [Extract] ¦ Confidence: 50%
Note on G3: [Extract] ¦ Confidence: 50%
Note on H3: [Extract] ¦ Confidence: 50%
Note on I3: [Extract] ¦ Confidence: 50%
Note on K3: [Extract] ¦ Confidence: 50%
Note on L3: [Extract] ¦ Confidence: 50%

## Instructions
UofA Excel Template — Instructions
What is this? | This workbook captures a credibility assessment for one Context of Use (COU). Fill it out, then run `uofa import` to convert it to a machine-verifiable evidence package.
One workbook = one COU | If your project has multiple COUs (e.g., different device classes or risk levels for the same model), create a separate workbook for each.
Profiles | Minimal (7 fields, lightweight audit trail) is for early-stage projects or live pipeline capture. Complete (full V&V 40 assessment) is for regulatory submissions.
Color coding | Blue cells = required for Minimal profile. Green cells = required for Complete profile. White cells = optional.
Credibility Factors | Sheet 4 is pre-populated with all 13 V&V 40 factors from Table 5-1. You do not need to assess all 13 — set Assessed = "No" or "N/A" for factors not relevant to your COU.
Dropdowns | Most constrained fields use dropdown lists. If you need a custom value not in the dropdown, type it directly (the validation will warn but not prevent entry).
After filling out | Run: uofa import your-file.xlsx --sign --key keys/your.key --check
What the tool checks | Gate 1 (SHACL): Are all required fields present? Gate 2 (Jena rules): Are there quality gaps (weakeners)? Gate 3 (Integrity): Is the hash and signature valid?
More info | https://crediblesimulation.com ¦ https://github.com/cloudronin/uofa

## _Lists (hidden)
Profile | Device Class | Model Risk Level | Assurance Level | Standards Reference | Entity Type | Has UQ | Pass/Fail | Factor Level (1-5) | Factor Status | Decision Outcome | Has UQ (COU-level)
Minimal | Class I | MRL 1 | Low | ASME-VV40-2018 | Requirement | Yes | Pass | 1 | assessed | Accepted | Yes
Complete | Class II | MRL 2 | Medium | NASA-STD-7009B | Model | No | Fail | 2 | not-assessed | Not accepted | No
| Class III | MRL 3 | High | FDA-2023-CMS | Dataset |  | Inconclusive | 3 | scoped-out | Conditional
| Category A | MRL 4 |  | Custom |  |  | N/A | 4 | not-applicable
| Category B | MRL 5 | [5 empty] | 5
Category C
Category D
Category E
Other

## Model & Data
Model & Data — What you modeled and compared against
Entity Type | Name | Identifier / URI | Description | Version | Source
Requirement | Micromotion and Contact Screening Requirement | Stable URI or local ID | Model must predict interface micromotion with sufficient fidelity (validated error ≤ ~10%, 95th-percentile micromotion < 50 µm osseointegration threshold) and contact footprint within acceptable tolerance to support design-freeze and cadaveric test go/no-go decision for Alpha-52 cup. | Version number (optional) | Where this came from (optional)
Model | ANSYS Mechanical FEA Model — Alpha-52 Acetabular Cup |  | TET10 finite element model of 52 mm porous Ti-6Al-4V shell seated in pelvis CT-derived geometry; Augmented Lagrange contact, Newton–Raphson solver, Latin hypercube UQ over key inputs.
Dataset | Foam Hemipelvis and Cadaveric Validation Dataset |  | Six 20 pcf foam hemipelvis blocks and two cadaveric pelves (T-scores −1.1, −2.0); LVDT and DIC displacement measurements at six rim points; Fuji film contact imprints under 1.5 kN and 30 N·m torque loading.
Dataset | Material and Geometry Input Dataset |  | Ti-6Al-4V properties per ASTM F136 (E = 110 GPa, ν = 0.34); bone E(ρ) from Hounsfield-to-modulus (Keyak); seating misalignment ±2° from build/jig tolerances; initial gap normal(50 µm, σ = 30 µm); calibrated friction coefficient μ = 0.56 ± 0.08 from three independent push-out tests.
Note on A3: [Extract] ¦ Confidence: 50%
Note on B3: [Extract] ¦ Confidence: 50%
Note on D3: [Extract] ¦ Confidence: 50%
Note on A4: [Extract] ¦ Confidence: 50%
Note on B4: [Extract] ¦ Confidence: 50%
Note on D4: [Extract] ¦ Confidence: 50%
Note on A5: [Extract] ¦ Confidence: 50%
Note on B5: [Extract] ¦ Confidence: 50%
Note on D5: [Extract] ¦ Confidence: 50%
Note on A6: [Extract] ¦ Confidence: 50%
Note on B6: [Extract] ¦ Confidence: 50%
Note on D6: [Extract] ¦ Confidence: 50%

## Validation Results
Validation Results — What you found when you tested the model
Result Name | Identifier / URI | Description | Compares To | Has UQ? | UQ Method | Metric Value | Pass / Fail
Interface Slip Regression vs. Experimental LVDT/DIC | ValidationResult | What was done, what was found | Model-predicted interface slip compared against LVDT and DIC measurements at six rim points across foam and cadaveric test articles under 1.5 kN compressive load plus 30 N·m torque. | LVDT and DIC measurements on foam hemipelvis and cadaveric pelves | Yes | Latin hypercube sampling (N=300) over friction coefficient, bone modulus scaling, seating angle, and initial gap | Regression slope 0.95, R² = 0.93, mean absolute percent error 8% (overall), 9% (cadaveric only); 95th-percentile micromotion 42 µm vs. 50 µm threshold | Pass
Contact Footprint Comparison vs. Fuji Film Imprints | ValidationResult |  | Predicted contact area and centroid location compared against Fuji film pressure-sensitive imprints under identical loading conditions. | Fuji film contact imprints | No |  | Contact area within 12%; centroid location within 3 mm | Pass
Mesh Convergence and GCI Study | ValidationResult |  | Three-mesh refinement study (0.9M, 1.3M, 1.8M elements) to assess discretization error; Richardson extrapolation applied; asymptotic convergence verified. | Richardson extrapolation / asymptotic convergence criterion | No | Grid Convergence Index (GCI) | GCI ≈ 2% for peak interface slip; GCI ≈ 5% for peak von Mises; asymptotic slope ≈ 2.1 | Pass
Numerical Solver Verification (Analytical Benchmarks) | ValidationResult |  | Vendor patch tests and in-house benchmarks (cantilever beam, Hertz contact) run against analytical solutions to verify code correctness. | Analytical solutions for cantilever displacement and Hertz contact pressure | No |  | Displacement error 0.7–1.3%; pressure error within same range | Pass
Note on A3: [Extract] ¦ Confidence: 50%
Note on B3: [Extract] ¦ Confidence: 50%
Note on D3: [Extract] ¦ Confidence: 50%
Note on E3: [Extract] ¦ Confidence: 50%
Note on F3: [Extract] ¦ Confidence: 50%
Note on G3: [Extract] ¦ Confidence: 50%
Note on H3: [Extract] ¦ Confidence: 50%
Note on I3: [Extract] ¦ Confidence: 50%
Note on A4: [Extract] ¦ Confidence: 50%
Note on B4: [Extract] ¦ Confidence: 50%
Note on D4: [Extract] ¦ Confidence: 50%
Note on E4: [Extract] ¦ Confidence: 50%
Note on F4: [Extract] ¦ Confidence: 50%
Note on H4: [Extract] ¦ Confidence: 50%
Note on I4: [Extract] ¦ Confidence: 50%
Note on A5: [Extract] ¦ Confidence: 50%
Note on B5: [Extract] ¦ Confidence: 50%
Note on D5: [Extract] ¦ Confidence: 50%
Note on E5: [Extract] ¦ Confidence: 50%
Note on F5: [Extract] ¦ Confidence: 50%
Note on G5: [Extract] ¦ Confidence: 50%
Note on H5: [Extract] ¦ Confidence: 50%
Note on I5: [Extract] ¦ Confidence: 50%
Note on A6: [Extract] ¦ Confidence: 50%
Note on B6: [Extract] ¦ Confidence: 50%
Note on D6: [Extract] ¦ Confidence: 50%
Note on E6: [Extract] ¦ Confidence: 50%
Note on F6: [Extract] ¦ Confidence: 50%
Note on H6: [Extract] ¦ Confidence: 50%
Note on I6: [Extract] ¦ Confidence: 50%

## Credibility Factors
Credibility Factors — V&V 40 Table 5-1 Assessment
Complete profile only. Assess each relevant factor: set Required Level (goal given model risk) and Achieved Level (what evidence supports).
Factor Type | Category | Required Level | Achieved Level | Acceptance Criteria | Rationale | Factor Status | Linked Evidence
V&V 40 factor (pre-filled, do not edit) | Grouping (pre-filled) | Credibility goal for this factor (1-5) | Level achieved by evidence (1-5) | Goal and basis for the required level (from V&V 40 Tables 3-4) | Brief justification (esp. if achieved < required) | assessed / not-assessed / scoped-out / not-applicable | URI of validation result that supports this factor (optional)
Software quality assurance | Verification — Code | 3 | 4 | Commercial solver with documented verification; in-house scripts under version control and CI testing | ANSYS Mechanical 2023 R2 (commercial solver with vendor patch tests passed). In-house Python preprocessors have 86% unit test coverage with CI on Jenkins. All simulation runs tracked under Git tag sim-hip-cup@v1.7.2 with fixed seeds and logged workstation specs. Requirements-to-model-to-test traceability matrix maintained in Confluence (CQ-AC-052); input decks, meshes, and scripts archived and hash-checked. | assessed
Numerical code verification | Verification — Code | 3 | 4 | Analytical solution comparison within acceptable tolerance for displacement and contact pressure | Vendor patch tests confirmed passed. In-house benchmark suite includes cantilever beam (displacement error 0.7–1.3%) and Hertz contact (pressure error 0.7–1.3%) compared to analytical solutions. Results documented in the simulation report. | assessed
Discretization error | Verification — Calculation | 3 | 4 | GCI < 5% for primary QoI (interface slip); asymptotic convergence confirmed | Three-mesh study (0.9M, 1.3M, 1.8M elements); asymptotic slope ≈ 2.1 confirms convergence. GCI ≈ 2% for peak interface slip (primary QoI) and 5% for peak von Mises. Final runs performed on finest mesh (1.8M elements). Contact penalty sensitivity ±25% causes < 3% change in micromotion. | assessed
Numerical solver error | Verification — Calculation | 3 | 4 | Residual tolerance documented; global reaction balance check performed; iterative convergence confirmed | Newton–Raphson with line search; force residual tolerance 1e-6; contact penetration limit 3 µm; automatic stabilization disabled. Last two equilibrium iterations show strain energy change < 0.1%. Global reactions match applied loads within 0.5%. Reaction balance check added following external red-team review. | assessed
Use error | Verification — Calculation | 3 | 4 | Independent review of model setup; boundary condition and mesh quality verification; process checklist signed | External SME (L. Gomez, OrthoSim LLC) performed red-team readout; two minor actions identified and closed (contact stabilization turned off; reaction balance check added). Team trained on ANSYS contact best practices; process checklist signed. Post-processing sanity checks performed: no stress oscillations at contact edge; spot checks with path integration agree within 4%. | assessed
Model form | Validation — Model | 3 | 3 | Governing equations and constitutive choices justified; known limitations documented | Linear elastic assumption for bone and cup justified by literature supporting elastic micromotion thresholds for osseointegration (< 50 µm). Frictional sliding with Augmented Lagrange contact and large-deflection enabled. Porous coating approximated by effective friction coefficient and measured compliance rather than geometric resolution — acknowledged limitation. No remodeling or press-fit plasticity modeled. Severe osteoporosis, revision cases, cemented shells, long-term remodeling, and extreme torque (> 40 N·m) explicitly excluded and documented in use restrictions. | assessed
Model inputs | Validation — Model | 3 | 4 | Material properties from standards/testing; boundary conditions from measurements; input uncertainties characterized | Ti-6Al-4V properties from ASTM F136 (E = 110 GPa, ν = 0.34). Bone E(ρ) from CT Hounsfield units via Keyak mapping with density cutoffs. Seating misalignment ±2° based on build/jig tolerance measurements. Initial gap characterized as normal(50 µm, σ = 30 µm). Friction coefficient calibrated on three independent push-out tests (μ = 0.56 ± 0.08); holdout specimens not used for calibration and were reserved for validation. Sobol sensitivity analysis identifies μ (0.44) and bone modulus (0.31) as dominant input drivers. | assessed
Test samples | Validation — Comparator | 3 | 3 | Multiple specimens including biofidelic material; statistical characterization of sample properties | Six foam hemipelvis blocks (20 pcf) and two cadaveric pelves (T-scores −1.1 and −2.0) used. Foam provides repeatable controlled baseline; cadaveric specimens provide biofidelic response. T-scores documented. Sample size is modest (especially cadaveric n=2) but fixtures verified and load paths confirmed by instrumented potting. Bone density characterization via T-score rather than full statistical distribution. | assessed
Test conditions | Validation — Comparator | 3 | 3 | Controlled and measured loading; calibrated instrumentation; boundary conditions verified | Loading applied via hip simulator at 1.5 kN compressive plus 30 N·m torque representing worst-case stance phase. Displacement measured at six rim points via LVDTs and DIC. Fixtures designed to mimic boundary stiffness; load paths verified by instrumented potting. Measurement uncertainty not explicitly quantified in the report beyond instrument identification; calibration status of LVDT/DIC not explicitly stated. | assessed
Equivalency of input parameters | Validation — Assessment | 3 | 3 | Model boundary conditions match test conditions; input parameter comparison documented | Model loading conditions (1.5 kN + 30 N·m) directly replicate hip simulator settings. Seating geometry and misalignment aligned with experimental fixture tolerances. Friction coefficient calibrated on independent push-out tests and reused without retuning for validation articles, demonstrating consistent input application. Initial gap distribution derived from build tolerances applied consistently. Full uncertainty propagation performed in UQ study matching physical input ranges. | assessed
Output comparison | Validation — Assessment | 3 | 4 | Quantitative error metrics for primary QoI; multiple comparison points; uncertainty bands reported | Interface slip comparison: regression slope 0.95, R² = 0.93, mean absolute percent error 8% overall and 9% cadaveric-only across six rim measurement points. Contact footprint: area within 12%, centroid within 3 mm vs. Fuji film imprints. LHS UQ (N=300) yields micromotion = 28 ± 7 µm at critical node with 95th percentile 42 µm, below 50 µm threshold. Residuals confirmed unbiased. Multiple metrics (slip magnitude, contact area, centroid) used. | assessed
Relevance of the quantities of interest | Applicability | 3 | 4 | QoI directly reflects osseointegration risk and contact mechanics; measurable both computationally and experimentally | Primary QoI is interface micromotion (µm), directly linked to osseointegration threshold (< 50 µm per literature). Secondary QoI is contact footprint geometry, relevant to load transfer assessment. Both quantities are directly measurable experimentally (LVDT, DIC, Fuji film) and computationally. QoI selection explicitly tied to the go/no-go decision criteria. Peak von Mises stress retained as supplementary QoI for screening but acknowledged as less decision-driving. | assessed
Relevance of the validation activities to the COU | Applicability | 3 | 3 | Validation conditions cover the intended operating envelope; gaps documented | Validation covers foam and osteopenic cadaveric specimens (T-score −1.1 to −2.0) matching the stated COU bone quality range. Loading (1.5 kN, 30 N·m) matches worst-case gait scenario in COU. Shell size 52 mm tested; applicability stated for 48–56 mm by interpolation. Known gaps explicitly documented: severe osteoporosis (T < −2.5), revision cases, cemented shells, remodeling, and torque > 40 N·m are out of scope. One osteoporotic cadaver specimen proposed as add-on to probe low-E tail. Cadaveric sample size (n=2) limits statistical coverage of bone quality variability within the covered range. | assessed
Note on C5: [Extract] ¦ Confidence: 50%
Note on D5: [Extract] ¦ Confidence: 50%
Note on E5: [Extract] ¦ Confidence: 50%
Note on F5: [Extract] ¦ Confidence: 50%
Note on G5: [Extract] ¦ Confidence: 50%
Note on C6: [Extract] ¦ Confidence: 50%
Note on D6: [Extract] ¦ Confidence: 50%
Note on E6: [Extract] ¦ Confidence: 50%
Note on F6: [Extract] ¦ Confidence: 50%
Note on G6: [Extract] ¦ Confidence: 50%
Note on C7: [Extract] ¦ Confidence: 50%
Note on D7: [Extract] ¦ Confidence: 50%
Note on E7: [Extract] ¦ Confidence: 50%
Note on F7: [Extract] ¦ Confidence: 50%
Note on G7: [Extract] ¦ Confidence: 50%
Note on C8: [Extract] ¦ Confidence: 50%
Note on D8: [Extract] ¦ Confidence: 50%
Note on E8: [Extract] ¦ Confidence: 50%
Note on F8: [Extract] ¦ Confidence: 50%
Note on G8: [Extract] ¦ Confidence: 50%
Note on C9: [Extract] ¦ Confidence: 50%
Note on D9: [Extract] ¦ Confidence: 50%
Note on E9: [Extract] ¦ Confidence: 50%
Note on F9: [Extract] ¦ Confidence: 50%
Note on G9: [Extract] ¦ Confidence: 50%
Note on C10: [Extract] ¦ Confidence: 50%
Note on D10: [Extract] ¦ Confidence: 50%
Note on E10: [Extract] ¦ Confidence: 50%
Note on F10: [Extract] ¦ Confidence: 50%
Note on G10: [Extract] ¦ Confidence: 50%
Note on C11: [Extract] ¦ Confidence: 50%
Note on D11: [Extract] ¦ Confidence: 50%
Note on E11: [Extract] ¦ Confidence: 50%
Note on F11: [Extract] ¦ Confidence: 50%
Note on G11: [Extract] ¦ Confidence: 50%
Note on C12: [Extract] ¦ Confidence: 50%
Note on D12: [Extract] ¦ Confidence: 50%
Note on E12: [Extract] ¦ Confidence: 50%
Note on F12: [Extract] ¦ Confidence: 50%
Note on G12: [Extract] ¦ Confidence: 50%
Note on C13: [Extract] ¦ Confidence: 50%
Note on D13: [Extract] ¦ Confidence: 50%
Note on E13: [Extract] ¦ Confidence: 50%
Note on F13: [Extract] ¦ Confidence: 50%
Note on G13: [Extract] ¦ Confidence: 50%
Note on C14: [Extract] ¦ Confidence: 50%
Note on D14: [Extract] ¦ Confidence: 50%
Note on E14: [Extract] ¦ Confidence: 50%
Note on F14: [Extract] ¦ Confidence: 50%
Note on G14: [Extract] ¦ Confidence: 50%
Note on C15: [Extract] ¦ Confidence: 50%
Note on D15: [Extract] ¦ Confidence: 50%
Note on E15: [Extract] ¦ Confidence: 50%
Note on F15: [Extract] ¦ Confidence: 50%
Note on G15: [Extract] ¦ Confidence: 50%
Note on C16: [Extract] ¦ Confidence: 50%
Note on D16: [Extract] ¦ Confidence: 50%
Note on E16: [Extract] ¦ Confidence: 50%
Note on F16: [Extract] ¦ Confidence: 50%
Note on G16: [Extract] ¦ Confidence: 50%
Note on C17: [Extract] ¦ Confidence: 50%
Note on D17: [Extract] ¦ Confidence: 50%
Note on E17: [Extract] ¦ Confidence: 50%
Note on F17: [Extract] ¦ Confidence: 50%
Note on G17: [Extract] ¦ Confidence: 50%

## Decision
Decision — The credibility judgment
Decision Outcome | Decision Rationale | Criteria Set | Decided By | Decision Date
Accepted | The Alpha-52 structural model meets the stated credibility targets for the intended decision. The primary QoI (interface micromotion) is validated with mean absolute percent error of 8–9%, GCI of ~2%, and a UQ-based 95th-percentile prediction of 42 µm against a 50 µm threshold, providing adequate margin. Contact footprint agreement is within 12% area and 3 mm centroid. Numerical uncertainty (GCI ~2%) is small relative to physical uncertainty. All open actions from independent red-team review are closed. Known applicability limits (severe osteoporosis, revision, extreme torque) are explicitly documented. The model is accepted for design-freeze support and progression to cadaveric testing within the stated applicability envelope. Pending items (alternative coating friction data, orthotropic bone sensitivity) are noted as low-risk add-on tasks that do not affect the current decision. | Auto-filled from Assessment Summary | R. Patel, Simulation Group | 2023-01-01
Note on A3: [Extract] ¦ Confidence: 50%
Note on B3: [Extract] ¦ Confidence: 50%
Note on D3: [Extract] ¦ Confidence: 50%
Note on E3: [Extract] ¦ Confidence: 50%